## Slide 1: A compact injection-locked UV laser system for quantum information processing
Nikolay Videnov
Quantum Information with Trapped Ions (QITI) Group

## Slide 2: Quantum Information With Trapped Ions
New lab
Quantum simulator using trapped ions
Tackle computationally hard problems using an engineered Hamiltonian

## Slide 3: UV lasers
Uses
  Laser Cooling
  Optical pumping
  State Detection
Getting UV light
  Frequency doubling Ti:Sapph
  External Cavity Diode Laser
  Injection locking

## Slide 4: Cavity Laser
Cavity laser
  Gain medium electronically pumped
  Saturation occurs forcing lasing
Injection locking
  Inject seed signal
  Above threshold power is diverted into injected singal

## Slide 5: Injection Locking
Cavity laser
  Gain medium electronically pumped
  Saturation occurs forcing lasing
Injection locking
  Inject seed signal
  Above threshold power is diverted into injected singal

## Slide 6: Physically realizing Injection Locking
Need mW-µW from a seed laser
Faraday isolator
Slave laser diode
T. Hosoya et al., Review of Scientific Instruments 86, 2015

## Slide 7: Diode Selection
Locking only happens around the free running oscillation frequency of the diode.
Need a laser diode around 369nm
The closest commercially available diode can be frequency selected to ~371nm*
  Cool to 369nm?
*Nichia NDU4116 (370-380nm)

## Slide 8: Test Fixture
20°C below lab temperature
  Achievable with Thermoelectric Coolers (TEC)
A few heat load calculations later we pick a pair of TECs
  Make a test block
  Add a dummy load
Dummy load
Test diode holder

## Slide 9: Laser Enclosure
Cooling below the dew point causes
  Condensation
  frost
Pressurize with dry nitrogen
Feedthroughs
  Optical
  Pressurized air
  electrical
Dummy load
Test diode holder
TEC
Water Cooling
block
Water Cooling
block
TEC
Frost

## Slide 11: Laser Enclosure
Machined an enclosure
Time to test!
¼” NPT ball valve
O-ring
Water cooling
bock
Diode Holder
D-sub connector
Pressure Sensor

## Slide 12: Climate Measurement
There’s a leak!
Make a blank and seal the D-sub connector

## Slide 13: A fix all solution
5 minute epoxy
A measurable improvement!

## Slide 14: Putting it all together
Machined diode holder and enclosure
Compact collimation
Water-cooling
Constant current controller

## Slide 15: Experimental setup

## Slide 16: Results
Cooled diode without condensation

## Slide 17: Conclusion/Outlook
Injection locking is an excellent way to amplify laser power
The challenges of injection locking can be tackled effectively with this design

## Slide 18
Quantum Information with Trapped Ions (QITI)
research.iqc.uwaterloo.ca/qiti/
Chung-You (Gilbert) Shih
Nikhil Kotibhaskar
Sainath Motlakunta
G R A D U A T E S T U D E N T S
Roland Hablützel
P O S T D O C S
Fereshteh Rajabi
U N D E R G R A D S
Yi Hong
Teoh
Rajibul
Islam
P I

## Slide 19: Injection Locking

## Slide 22: Head Load

## Slide 23: UV lasers
Uses
  Micro lithography
  Laser ablation
  Trapped ions
Getting UV light
  Frequency doubling Ti:Sapph
  External Cavity Diode Laser
  Injection locking